## For Web (hidden)
|  |  | 1a | 1b | 2 | 3 | 4 | 5 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
|  | Specialty | # cases, irrig | # cases, non-irrig | Use mono but should use bip | Footpad diff to locate | Ever std back from OR table to look for foot pedal during surgery? | # of times need help moving foot pedal | Ask asst to active foot pedal | Is distraction/time delay finding foot pedal a Potential Safety/Qual Issue? | Time delay caused unnec blood loss | Ever Pressed wrong foot pedal during sgy | Is bipolar foot pedal the only foot pedal under the table? | Balance oK? | Does BiPAD interfere with ability to see surgical site? | BiPAD switch adequately sensitive? | Will BiPAD improve use of bipolar forceps? | Would you alternate between the foot pedal and the hand switch? | Will you need to remove the actuator assembly in some situations | Does the prototype fit comfortably in your hands | Would you still use the foot pedal instead o the hand switch | Does the finger guard help avoid accidental activation? | Do you find it easier to activate the hand switch or Foot pedal>
|  |  | # | # | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | # | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=good, 0=challenging | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Hand sw, 0=foot pedal
1 |  | Neurosurgeon | 0 | 10 | 1 | 1 | 1 | 4 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1
2 |  | Neurosurgeon | 0 | 10 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
3 |  | Plastics | 3 | 0 | 0 | 1 | 1 | 4 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | Not sure | Not sure | Not sure | 1 | Not sure | Not sure | 1
4 |  | Neurosurgeon | 6 | 6 | 0 | 1 | 1 | 3 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0
5 |  | Plastics | 0 | 3 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1
6 |  | Plastics | 2 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1
7 |  | Neurosurgeon | 7 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0
8 |  | Neurosurgeon | 6 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1
9 |  | Neurosurgeon | 0 | 3 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | - | 1
10 |  | Neurosurgeon | 2 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
11 |  | Neurosurgeon | 5 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0
12 |  | Neurosurgeon | 7 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0
13 |  | Neurosurgeon | 0 | 6 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0
14 |  | Neurosurgeon | 4 | 8 | 1 | 1 | 1 | 3 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0
15 |  | Neurosurgeon | 1 | 1 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1
16 |  | Neurosurgeon | 5 | 5 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1
17 |  | Neurosurgeon | 2 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
18 |  | Neurosurgeon | 3 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1
| Totals: | 3 Plastic surgeons | 53 | 55 | 10 | 17 | 17 | 36 | 10 | 14 | 8 | 17 | 2 | 14 | 0 | 16 | 14 | 8 | 10 | 17 | 7 | 11 | 12
| # responded to this question |  | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 17 | 17 | 17 | 18 | 17 | 16 | 18
|  | 15 Neorosurgeon | 2.9444 | 2.92 | 55.56% | 94.44% | 94.44% | 200% | 55.56% | 77.78% | 44.44% | 94.44% | 11.11% | 77.78% | 0% | 88.89% | 82.35% | 47.06% | 58.82% | 94.44% | 41.18% | 68.75% | 66.67%

## Original
|  |  | 1a | 1b | 2 | 3 | 4 | 5 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | 19 | 20 | 21
|  | Specialty | # casss, irrig | # cases, non-irrig | Use mono but should use bip | Footped diff to locate | Ever std back from OR table to look for foot pedal during surgery? | # of times need help moving foot pedal | Ask asst to active foot pedal | Is distraction/time delay finding foot pedal a Potential Safety/Qual Issue? | Time delay caused unnec blood loss | Ever Pressed wrong foot pedal during sgy | Is bipolar foot pedal the only foot pedal under the table? | Balance oK? | Does Bipad interfere with ability to see surgical site? | BiPAD switch adequately sensitive? | Will BiPAD improve use of bipolar forceps? | Would you alternate between the foot pedal and the hand switch? | Will you need to remove the actuator assembly in some situations | Does the prototype fit comfortably in your hands | Would you still use the foot pedal instead o the hand switch | Does the finger guard help avoid accidental activation? | Do you find it easier to activate the the hand switch or Foot pedal>
|  |  | # | # | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | # | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=good, 0=challenging | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Y, 0=N | 1=Hand sw, 0=footpedal
1 | RW | Nsgn | 0 | 10 | 1 | 1 | 1 | 4 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1
2 | BC | Nsgn | 0 | 10 | 1 | 1 | 1 | 2 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
3 | OT | plastics | 3 | 0 | 0 | 1 | 1 | 4 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | Not sure | Not sure | Not sure | 1 | Not sure | Not sure | 1
4 | TS | Nsgn | 6 | 6 | 0 | 1 | 1 | 3 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0
5 | JS | plastics | 0 | 3 | 0 | 1 | 1 | 2 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 1
6 | HS | plastics | 2 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1
7 | EE | Nsgn | 7 | 0 | 0 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 0 | 0
8 | MS | Nsgn | 6 | 0 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1
9 | JA | Nsgn | 0 | 3 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | - | 1
10 | AM | Nsgn | 2 | 0 | 0 | 1 | 1 | 2 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
11 | KK | Nsgn | 5 | 0 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 0 | 1 | 1 | 1 | 0
12 | KW | Nsgn | 7 | 0 | 1 | 0 | 0 | 2 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0
13 | SR | Nsgn | 0 | 6 | 0 | 1 | 1 | 2 | 0 | 0 | 0 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 1 | 1 | 0 | 0
14 | KS | Nsgn | 4 | 8 | 1 | 1 | 1 | 3 | 1 | 1 | 0 | 1 | 0 | 0 | 0 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0
15 | RN | Nsgn | 1 | 1 | 1 | 1 | 1 | 3 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 1 | 1 | 0 | 1 | 1
16 | DS | Nsgn | 5 | 5 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 0 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 0 | 1
17 | DT | Nsgn | 2 | 3 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1 | 0 | 1 | 1
18 | T | Nsgn | 3 | 0 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 0 | 1 | 0 | 0 | 1 | 0 | 1 | 1 | 0 | 0 | 1
|  | 3 plastic surgeons | 53 | 55 | 10 | 17 | 17 | 36 | 10 | 14 | 8 | 17 | 2 | 14 | 0 | 16 | 14 | 8 | 10 | 17 | 7 | 11 | 12
|  |  | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 18 | 17 | 17 | 17 | 18 | 17 | 16 | 18
|  | 15 Nsgn | 2.9444 | 2.92 | 55.56% | 94.44% | 94.44% | 200% | 55.56% | 77.78% | 44.44% | 94.44% | 11.11% | 77.78% | 0% | 88.89% | 82.35% | 47.06% | 58.82% | 94.44% | 41.18% | 68.75% | 66.67%
Totals:
# responded to this question
1. 53 cases per week using irrigating, 55 cases per week using non-irrigating
2.Ten (10) surgeons said they used monopolar when they should use bipolar, (generally because its easier, has hand switch)
3. 17 of 18 ( 94 %) find it difficulty to locate the bipolar foot pedal during surgery.
4. 17 of 18 ( 94 %) stand back from the table during surgery to find the foot pedal.
5. Surgeons reported that staff need to crawl on the floor to locate a foot pedal about 2x per case.
6. No question.
7. 10 of 18 (56 %) surgeons ask an assistant to press the foot pedal for them.
8. 14 of 18 (78 %) surgeons said that the time delay/distsraction of finding the foot peal is a safety/quality issue.
9. 8 of 18 ( 44 % )surgeons said that time delay associated with finding a foot pedal caused unnecessary blood loss.
10. 17 of 18 ( 94 %) surgeons said they have pressed the wrong foot pedal during surgery.
11. 16 of 18 ( 89 %) surgeons said that the bipolar foot pedal is NOT the only foot pedal under the table.
12. 14 of 18 ( 78%) surgeons said that the balance of the BiPAD cord/hand switch was adequate.
13. 0 of 18 ( Zero %) surgeons said that the BiPAD hand switch interferes with ability to see the surgical site.
14. 16 of 18 (89 %) said that the BiPAD hand switch was adequately sensitive.
15. 14 of 17 ( 82 %) surgeons said that the BiPAD hand switch will improve the use of bipolar forceps.
16. 8 of 17 ( 47 %) surgeons said that they would alternate between the BipAD and the foot pedal.
17. 10 of 17 ( 59 %) surgeons said they might need to remove the BiPAD hand switch in some situations.
18. 17 of 18 (94 %) surgeons said the that BiPADS fits comfortably in their hands.
19. 7 of 17 ( 41 %) surgeons said that they would still use the foot pedal instead of the hand switch,
20. 11 of 16 ( 69 %) surgeons said that the finger guard helps reduce the risk of accidental activation.
21. 12 of 18 ( 67 %) surgeons said that they find it easier to activate the bipolar with the hand switch than the foot pedal. 3 of 3 (100 %) plastic surgeons said the hand switch was easier to activate.
BiPAD® Surgical, Copyright, 2019